DEPARTMENT OF STATE POLICE

TRAFFIC SAFETY DIVISION TRAINING DIVISION

TESTS FOR BREATH ALCOHOL

Filed with the secretary of state on

These rules become effective take effect immediately after upon filing with the secretary of state unless adopted under section 33, 44, or 45a(9) of the administrative procedures act of 1969, 1969 PA 306, MCL 24.233, 24.244, or 24.245a.  Rules adopted under these sections become effective 7 days after filing with the secretary of state.

(By authority conferred on the department of state police by section 190 of the Aeronautics code of the state of Michigan, 1945 PA 327, MCL 259.190, and section 625a of the Michigan vehicle code, 1949 PA 300, MCL 257.625a)

R 335.2651, R 325.2652, R 325.2653, R 325.2655, and R 325.2658 of the Michigan Administrative Code are amended, R 335.2656a, R 335.2657a, and R 335.2659 are added, and R 335.2654 is rescinded, as follows:

R 325.2651  Definitions.

Rule 1.  (1) As used in these rules:

(a) "Acts" means section 190 of the Aeronautics code of the state of Michigan, 1945 PA 327, MCL 259.190; section 625a of the Michigan vehicle code, 1949 PA 300, MCL 257.625a; section 359 of the railroad code of 1993, 1993 PA 354, MCL 462.359; and sections 80181, 81136, and 82137 of the natural resources and environmental protection act, 1994 PA 451, MCL 324.80181, 324.81136, and 324.82137. means 1945 PA 327, MCL 259.190, 1949 PA 300, MCL 257.625a, 1993 PA 354, MCL 462.359, 1994 PA 451, MCL 324.80181, MCL 324.81136, and MCL 324.82137.

(b) "Alcohol standard" means a certified alcohol standard.

(c) "Calendar month" means 12:01 a.m. on the first date of any of the 12 named months of the year to midnight on the last date of the same month.

(d)(c) "Calendar week" means 12:01 a.m. Sunday to midnight Saturday.

(e)(d) "Class" means a classification of operator status as certified by the department, based on training and function as specified in R 325.2658. “Certified” means the operator completed the required training and possesses a certificate of training.

(f)(e) "Department" means the department of state police. "Class" means a classification of operator status as certified by the department, based on training and function as specified in R 325.2658.

(g)(f) "Equipment" means evidential and preliminary breath alcohol test instruments, simulator devices, calibration stations, forms, and any accessories and supplies necessary for compliance with the procedures in these rules or law. “Class I operator” is an operator certified to conduct a subject test with a preliminary breath testing instrument.

(h)(g) "Evidential breath alcohol analysis" means chemical analysis of an essentially alveolar breath sample that indicates a specific result in grams of alcohol per 210 liters of breath. “Class II operator” is an operator certified to calibrate a preliminary breath testing instrument.

(i)(h) "Evidential breath alcohol test instrument" means an evidential breath testing device that indicates a specific result in grams of alcohol per 210 liters of breath. “Class III operator” is an operator certified to conduct a subject test with an evidential breath testing instrument.

(j)(i) "Preliminary breath alcohol analysis" means chemical analysis of essentially alveolar breath samples that indicates the presence or absence of alcohol in a person's blood. “Class IVA operator” is an operator certified to instruct the class I, II, and III certification courses.

(k)(j) "Preliminary breath alcohol test instrument" means a breath alcohol screening device that indicates the presence or absence of alcohol in a person's blood. “Class IVB operator” is an operator certified to calibrate and repair an evidential breath testing instrument.

(f)(k)  "Department" means the department of state police.

(g)(l) "Equipment" means evidential and preliminary breath alcohol test instruments, simulator devices, calibration stations, forms, and any accessories and supplies necessary for compliance to comply with the procedures in these rules or law.

(h)(m) "Evidential breath alcohol analysis" means chemical analysis of an essentially alveolar breath sample that indicates a specific result in grams of alcohol per 210 liters of breath.

(i)(n) "Evidential breath alcohol test instrument" means an evidential breath testing device that indicates a specific result in grams of alcohol per 210 liters of breath.

(j)(o) "Preliminary breath alcohol analysis" means chemical analysis of essentially alveolar breath samples that indicates the presence or absence of alcohol in an individual’s person's blood.

(k)(p) "Preliminary breath alcohol test instrument" means a breath alcohol screening device that indicates the presence or absence of alcohol in an individual’s person's blood.

(l)(r) "Shall" means that a function is mandatory.

(m)(s) "Should" means that a function is recommended, but not mandatory, with reasonable deviation allowed.

(2) Terms used in the acts have the same meanings when used in these rules.

R 325.2652  Approved equipment.

Rule 2.  (1) Except as provided in subrule (2) of this rule,Evidential evidential and preliminary breath alcohol test instruments must shall be tested by, and shall meet the existing model specifications for evidential breath alcohol analysis as established by, the United States Department of Transportation, National Highway Traffic Safety Administration.  The specifications are identified as “Model Specifications for Devices to Measure Breath Alcohol” 49 FR 48855 et seq., (December 14, 1984), as amended by 58 FR 48705, (September 17, 1993), as amended by 82 FR 50940 et seq., (November 2, 2017) 62 FR 43416 et seq., (August 13, 1997), 69 FR 42237 et seq., (July 14, 2004), 71 FR 37159 et seq., (June 29, 2006), 72 FR 71480 et seq., (December 17, 2007), and FR Doc. 2012-14582 (June 14, 2012), and are adopted in these rules by reference.  A copy of the specifications and a current conforming products list are available from the United States U.S. Department of Transportation, National Highway Traffic Safety Administration, 1200 New Jersey Avenue SE, Washington, D.C. 20590, its website at nhtsa.gov,  or from the Michigan State Police, Training Division Records Resource Section, 7150 Harris Drive, P.O. Box 30634, Lansing, MI, 48909, its website at www.michigan.gov/msp/services/foia. A fee of $1.00 may be charged for each requested copy.  A copy of the specifications and conforming products list may be inspected at the Michigan State Police, Training Division, 7426 N Canal Road, Lansing, MI, 48913 during normal business hours.

(2) If the United States Department of Transportation discontinues the testing of evidential breath alcohol test instruments or the issuance of model specifications for that test such testing, only those instruments tested and approved by the department can shall be used.

(3) An application for equipment approval shall must be submitted to the department.

(4) Only Operators shall only use equipment approved by the department for evidential and preliminary breath alcohol analysis. shallbe utilized by operators.

R 325.2653  Equipment accuracy evidential breath test instruments.

Rule 3.  (1) For evidential breath test instruments that do not examine a known standard with each subject test, an An appropriate class operator who has been certified in accordance with R 325.2658 shall verify an evidential breath alcohol test instrument for accuracy at least once each calendar week, or more frequently as the department may require.  Alternatively, a pre-programmed self-test for accuracy may be initiated by the evidential breath alcohol test instrument and recorded by an appropriate class operator who has been certified in accordance with R 325.2658, if the instrument is so equipped and programmed.  The tests for accuracy are not required to be performed within 7 days of each other.  The test for accuracy must shall be made using an alcohol standard that is approved by the department.  For the evidential breath alcohol test instrument to meet the requirements for accuracy, a test result not exceeding +/- 5% must shall be obtained when using a controlled device. Controlled devices include both of the following:

(a) A Wwet bath device that delivers an alcohol vapor concentration test result of .080 grams of alcohol per 210 liters of vapor.

(b) A Ccompressed alcohol gas device that delivers a test result of .080 grams of alcohol per 210 liters of vapor before applying applicable barometric pressure correction.  To meet this requirement, the instrument must analyze the known ethanol gas standard within +/- 5% of the adjusted, or corrected alcohol concentration, based on the barometric correction.after applying applicable altitude or topographic elevation correction factor supplied by the manufacturer.  The correction factor may be applied by the operator or by the evidential breath alcohol test instrument if so pre-programmed.

(2) A weekly test is not required for evidential breath test instruments that examine a known alcohol standard, either wet bath or compressed gas, with each test subject.  If an accuracy check is conducted, then the results must be retained either in log form by the agency where the instrument is installed or electronically within the instruments memory.  An appropriate class operator who has been certified in accordance with R 325.2658 shall verify for accuracy a preliminary breath alcohol test instrument at least monthly, or more frequently as the department may require.  The test for accuracy shall be made using an alcohol standard that is approved by the department.  For the preliminary breath alcohol test instrument to meet the requirements for accuracy, a test result not exceeding +/- 5% shall be obtained when using a controlled device.  Controlled devices include both of the following:

(a) Wet bath device that delivers an alcohol vapor concentration of .080 grams of alcohol per 210 liters of vapor.

(b) Compressed alcohol gas device that delivers a test result of .080 grams of alcohol per 210 liters of vapor after applying applicable altitude or topographic elevation correction factor supplied by the manufacturer.  The correction factor may be applied by the operator or by the preliminary breath alcohol test instrument calibration station if so pre-programmed.

(3) Approved evidential breath alcohol test instruments that do not examine a known alcohol standard with each test subject must shall be inspected, verified for accuracy, and certified as to their proper working order within 120 days after of the previous inspection by either an appropriate class operator who has been certified in accordance with R 325.2658, or a manufacturer-trained representatives approved by the department.

(4) Approved evidential breath alcohol test instruments that examine a known alcohol standard with each subject test must be inspected, verified for accuracy, and certified as to their proper working order not less than 2 times annually by either an appropriate class operator who has been certified in accordance with R 325.2658, or a manufacturer-trained representative approved by the department.

R 325.2654 Rescinded. Equipment repair and service.

Rule 4. (l) The repair and service of equipment approved by the department for evidential and preliminary breath alcohol analysis shall be at the expense of the agency using the equipment.  The department or the agency using the equipment may arrange to have this service provided either by an appropriate class operator who has been certified in accordance with R 325.2658 or a manufacturer-trained authorized representatives approved by the department.

(2) After repair or service and before being placed in service, evidential and preliminary breath alcohol test instruments shall be verified for accuracy in accordance with the provisions of R 325.2653 and records of verification shall be kept as required by the department.

R 325.2655  Techniques and procedures – evidential breath test instruments.

Rule 5.  (1) A procedure that is used in conjunction with evidential breath alcohol analysis must shall be approved by the department and shall comply with the  be in compliance with all of the following requirements: provisions:

(a) Except as provided otherwise in these rules, evidential breath alcohol test instruments must shall only be operated by appropriate class operators who are have been certified in accordance with R 325.2658.

(b) All analyses must shall be conducted following procedures approved by the department and using forms approved by the department, as required.

(c) Records of operation, analyses, and results must shall be maintained at the evidential breath alcohol test instrument location as prescribed by the department, and copies must be copies forwarded to the department as required.

(d) The department shall test samples from each lot of alcohol standards used in this the state, in conjunction with evidential breath alcohol test instruments. The department shall certify for use those lots of alcohol standards that are found to be proper in chemical composition.

(e) An person individual may shall be administered an evidential breath alcohol analysis on an evidential breath alcohol test instrument only after being observed for not less than 15 minutes before collection of the breath sample by at least 1 appropriate class operators who that is  certified in accordance with R 325.2658. The observation period may be conducted by more than 1 operator working in concert.  During the observation period, the person individual shall not have smoked, regurgitated, or placed anything in his or her mouth, except for the mouthpiece associated with the performance of the evidential breath alcohol analysis.  The operator need not stare continuously at the person individual, but shall must be close enough to be aware of the individual’s person's actions and conditions.  The operator may complete paperwork, enter data into the evidential breath alcohol test instrument, or conduct other reasonable tasks during the observation period, provided if the individual person is within the operator's field of vision.  Breaks in the observation period lasting only a few seconds do not invalidate the observation  if the operator can reasonably determine that the individual person did not smoke, regurgitate, or place anything in his or her mouth during the breaks in the observation.

(f) The operator shall request a second evidential breath sample alcohol from the subject after the first sample is provided and analyzed by the instrument, unless an item person being tested and administered, unless the person refuses to give the second breath sample or a substance is found in the subject’s person's mouth after subsequent to the first evidential breath alcohol sample analysis that could interfere with the breath sample result.  Obtaining the first breath sample result is sufficient to meet the requirements for evidentiary purposes prescribed in section 625c of the Michigan vehicle code, 1949 PA 300, MCL 257.625c.  The purpose of obtaining a second breath sample result is to confirm the result of the first breath sample result.

(g) For instruments reporting 2 digits after the decimal, a A second breath sample result shall must not vary from the first breath sample result by more than the allowable variation listed in Table 1. Tables 1 and 2 read as follows:

Table 1

Breath alcohol allowable Allowable variation of second breath sample result range from the first breath sample result.

First Sample	Allowable variation

0.00 - 0.14	  	      +/- 0.01

0.15 - 0.24		      +/- 0.02

0.25 - 0.34		      +/- 0.03

0.35 - or more 	      +/- 0.04

For instruments reporting 3 digits after the decimal, the second breath sample result must not vary from the first sample result by more than the allowed variation listed in Table 2.

Table 2

Allowable variation of second breath sample result from the first breath sample result.

First Sample		Allowable Variation

0.000 - 0.149		+/- 0.010

0.150 – 0.249		+/- 0.020

0.250 – 0.349		+/- 0.030

0.350 or more		+/- 0.040

If the variation is more than that allowed in either table Table 1, the operator shall request a third breath sample from the subject person being tested and a third breath sample result may be obtained.  If the third breath sample result does not conform to the allowable variationvariations of either of the first 2 tests based on the allowable variation listed in either table, the subject  in Table 1 of either of the first 2 tests, the person shall be requested to submit a blood or urine sample for analysis in accordance with the acts MCL 257.625a and the procedures established in R 325.2671 to R 325.2677.

(gh) The results of an evidential breath alcohol analysis of a person's subject’s breath shall must be expressed in terms of grams of alcohol per 210 liters of breath, truncated to the second decimal place or expressed to 3 decimals if the instrument is programed to do so.  For example, 0.237 found shall be reported as 0.23.

(i) If the instrument analyzes a known ethanol standard during a subject’s breath test, the results of that analysis must be no lower than 0.074 g/210L and no higher than 0.084 g/210L of the nominal value of the standard.  If the analysis is not within the prescribed standard the instrument must abort the test indicating the ethanol standard was not within the required range.  The operator may attempt additional tests.

(2) A procedure that is used in conjunction with preliminary breath alcohol analysis must be approved by the department and shall be in compliance with all of the following provisions:

(a) Except as provided otherwise in these rules, preliminary breath alcohol test instruments shall only be operated by appropriate class operators who have been certified in accordance with R 325.2658

(b) A person may be administered a preliminary breath alcohol analysis on a preliminary breath alcohol test instrument only after the operator determines that the person has not smoked, regurgitated, or placed anything in his or her mouth for at least 15 minutes.

(c) Records shall be maintained at the preliminary breath alcohol test instrument location as prescribed by the department and copies shall be forwarded to the department as required.

(3) A person's welfare shall be protected by requesting medical assistance if the result of the evidential or preliminary breath alcohol analysis is 0.35 or more.  Failure to request medical assistance does not affect the validity or evidential value of the result of an evidential or preliminary breath alcohol analysis.

R 325.2656a  Equipment Accuracy – Preliminary Breath Test Instrument.

Rule 6a.  An appropriate class operator who has been certified in accordance with R 325.2658 shall verify for accuracy a preliminary breath alcohol test instrument at least monthly, or more frequently as the department may require.  The test for accuracy must be made using an alcohol standard that is approved by the department.  For the preliminary breath alcohol test instrument to meet the requirements for accuracy, a test result not exceeding +/- 5% must be obtained when using a controlled device.  Controlled devices include both of the following:

(a) A wet bath device that delivers an alcohol vapor concentration of .080 grams of alcohol per 210 liters of vapor.

(b) A compressed alcohol gas device that delivers a test result of .080 grams of alcohol per 210 liters of vapor before applying applicable altitude or topographic evaluation correction factor supplied by the manufacturer. The correction factor may be applied by the operator or by the preliminary breath alcohol test instrument calibration station, if pre-programmed.

R 325.2657a  Techniques and Procedures – Preliminary Breath Test Instruments.

Rule 7a.  (1) A procedure that is used in conjunction with preliminary breath alcohol analysis must be approved by the department and comply with all of the following:

(a) Except as provided otherwise in these rules, preliminary breath alcohol test instruments must only be operated by appropriate class operators who have been certified in accordance with R 325.2658.

(b) An individual  may be administered a preliminary breath alcohol analysis on a preliminary breath alcohol test instrument only after the operator determines that the individual has not smoked, regurgitated, or placed anything in his or her mouth for not less than 15 minutes.

(c) Records must be maintained at the preliminary breath alcohol test instrument location as prescribed by the department and copies must be forwarded to the department as required.

(2) An individual’s welfare must be protected by requesting medical assistance if the result of the evidential or preliminary breath alcohol analysis is 0.35 or more. Failure to request medical assistance does not affect the validity or evidential value of the result of an evidential or preliminary breath alcohol analysis.

R 325.2658  Operator training and certification.

Rule 8.  (1) The department shall train and certify personnel of law enforcement agencies to perform various functions as described in this rule, and shall designate those such individuals persons as class I, class II, class IIIA, class IIIB, class IVA, or class IVB operators.  An operator may hold multiple and concurrent classifications.  Operator certification is non-expiring.  To maintain a class III certification, each class III operator certified AFTER January 1, 2022 is required to re-certify every 2 years.  Class III operators certified BEFORE January 1, 2022  are required to recertify before December 31, 2023, and every 2 years after that.  Any class III operator who fails to re-certify before the current certification expires is required to attend and successfully complete the class III operator training as detailed in R 325.2658(2)(c).  An operator whose certification has lapsed can testify about incidents that occurred during their valid certification period.

(2) The minimum training requirements and proficiency standards for operator certification are as follows:

(a) A class I operator shall complete a class I training course approved by the department, and obtain a minimum score of 70% on a written examination administered by the department.  An individual that fails the examination is provided an opportunity to retake the written examination.  An individual that fails the second attempt to pass the written examination shall retake the class I training course and successfully pass the examination to qualify for the class I certification. And demonstrate proficiency in the operation of a preliminary breath alcohol test instrument to the department.  The department shall provide a person who fails to pass the written examination or demonstrate the required proficiency a second opportunity to take the written examination or demonstrate the required proficiency.  A person who fails on his or her second opportunity to successfully pass the written examination or demonstrate the required proficiency must retake the class I training course, successfully pass the written examination, and demonstrate the required proficiency, in order to qualify for class I operator certification.

(b) A class II operator shall complete a class II training course approved by the department, and obtain a minimum score of 70% on a written examination administered by the department.  An individual that fails the examination is provided an opportunity to retake the written examination.  An individual that fails the second attempt to pass the written examination shall retake the class I training course and successfully pass the examination to qualify for the class II certification. And demonstrate proficiency in the operation of an evidentiary breath alcohol test instrument to the department.  The department shall provide a person who fails to pass the written examination or demonstrate the required proficiency a second opportunity to take the written examination or demonstrate the required proficiency.  A person who fails on his or her second opportunity to successfully pass the written examination or demonstrate the required proficiency must retake the class II training course, successfully pass the written examination, and demonstrate the required proficiency, in order to qualify for class II operator certification.

(c) A class IIIA III operator shall be currently certified as a class I II operator, complete a class IIIA III training course approved by the department, obtain a minimum score of 70% on a written examination administered by the department, and demonstrate proficiency in the use of an evidential breath testing instrument.   verification and calibration of a preliminary breath alcohol test instruments to the department. An individual that fails the examination or fails to demonstrate the required proficiency is provided an opportunity to retake the written examination or demonstrate the required proficiency.  An  individual that fails the second attempt to pass the written examination or fails to demonstrate the required proficiency shall retake the class III training course and successfully pass the written examination and demonstrate the required proficiency to qualify for the class III certification. The department shall provide a person who fails to pass the written examination or demonstrate the required proficiency a second opportunity to take the written examination or demonstrate the required proficiency.  A person who fails on his or her second opportunity to successfully pass the written examination or demonstrate the required proficiency shall be required to retake the class IIIA training course, successfully pass the written examination, and demonstrate the required proficiency in order to qualify for class IIIIA operator certification.

(d) A class IIIB IVA operator shall be currently certified as a class II III operator, and complete a class IIIB IVA training course approved by the department., obtain a minimum score of 70% on a written examination administered by the department, and demonstrate proficiency in the limited service of an evidentiary breath alcohol test instruments to the department.  The department shall provide a person who fails to pass the written examination or demonstrate the required proficiency a second opportunity to take the written examination or demonstrate the required proficiency.  A person who fails on his or her second opportunity to successfully pass the written examination or demonstrate the required proficiency shall be required to retake the class IIIB training course, successfully pass the written examination, and demonstrate the required proficiency in order to qualify for class IIIB operator certification.

(e) A class IVA IVB operator shall be currently certified as a class III IVA operator.  Before class IVB certification, the class IVB operator shall receive additional training in the repair and service of evidential breath instruments from the manufacturer of the instruments or by a current class IVB operator, if the manufacturer is unavailable for training. complete a class IVA training course approved by the department, and demonstrate proficiency instructing class II and class IIIA operator training courses approved by the department.  The department shall provide a person who fails to demonstrate the required proficiency a second opportunity to demonstrate the required proficiency.  A person who fails on his or her second opportunity to demonstrate the required proficiency shall be required to retake the class IVA training course and demonstrate the required proficiency in order to qualify for class IVA operator certification.

(f) A class IVB operator shall be currently certified as a class IVA operator that has been designated by the department to administer the breath alcohol testing program for the state of Michigan.  Before class IVB designation, the class IVB operator shall receive additional training in the repair and service of evidential and preliminary breath alcohol test instruments by the manufacturer of such instruments.

(3) The department shall develop and distribute to each certified operator a training manual for each of the operator's classifications.  Training manuals must specify the functions performed by each class pursuant to this rule, as well as the knowledge and skills necessary to perform the appropriate functions.

(4) The primary functions of each class are described in Table 23.  Additional functions not described in Table 23 may be designated by the department and described and explained in the appropriate training manual.

(5) A person who An individual that meets the training requirements and proficiency standards for operator certification in accordance with this rule shall be is issued a certification card for the appropriate class by the department.  The certification card shall must remain the property of the department.

(6) A class I, class I II, class IIIA III, class IIIB, class IVA, and class IVB operator shall comply with all applicable department rules, policies, and procedures or the department may suspend his or her operator certification. A person whose operator certification has been suspended by the department shall return his or her certification card to the department within 7 days of receiving notice of the suspension.  An individual person shall may make a request for reinstatement of operator certification to the department in writing and, upon after approval, the department may require the individual person to complete an operator training class, pass a written examination, and/or demonstrate required proficiency, or any combination thereof.

(7) The requirements for each class of operators are included in Table 3 as follows:

Table 2 3

OPERATOR FUNCTION			OPERATOR CLASS

|  | I | II | IIIA | IIIB | IVA | IVB |
| --- | --- | --- | --- | --- | --- | --- |
| Train and certify class I operators |  |  | X |  | X | X |
| Train and certify class II operators |  |  |  |  | X | X |
| Train and certify class IIIA operators |  |  |  |  | X | X |
| Train and certify class IIIB operators |  |  |  |  |  | X |
| Train and certify class IVA operators |  |  |  |  |  | X |
| Administer preliminary breath alcohol analyses using preliminary breath alcohol test instruments approved for use by the department | X | X | X |  | X | X |
| Administer evidentiary breath alcohol analyses using evidential breath alcohol test instruments approved for use by the department |  | X |  | X | X | X |
| Verify for accuracy and calibrate preliminary breath alcohol test instruments approved for use by the department |  |  | X |  | X | X |
| Verify for accuracy evidential breath alcohol test instruments approved for use by the department |  | X |  | X | X | X |
| Inspect, certify, service, repair, and calibrate evidentiary breath alcohol test instruments approved for use by the department for proper working order |  |  |  |  |  | X |
| Provide service on preliminary breath alcohol test instruments approved for use by the department |  |  |  |  |  | X |
| Conduct inspections for compliance with applicable department rules, policies, and procedures |  |  |  |  |  | X |
| Provide limited service on evidentiary breath alcohol test instruments specified by the department |  |  |  | X |  |  |

|  | Class I | Class II | Class III | Class IVA | Class IVB |
| --- | --- | --- | --- | --- | --- |
| Administer preliminary breath alcohol analyses using preliminary breath alcohol test instruments approved for use by the department. | X | X | X | X | X |
| Calibrate preliminary breath alcohol test instruments approved for use by the department. |  | X | X | X | X |
| Administer evidential breath alcohol analyses using an evidential breath instrument approved by the department. |  |  | X | X | X |
| Record weekly verification for approved evidential breath test instruments that do not examine a known standard with each subject test. |  |  | X | X | X |
| Train and certify all lower-level class operators. |  |  |  | X | X |
| Calibrate and service an evidential breath testing instrument approved for use by the department. |  |  |  |  | X |

R 325.2659  Equipment Repair and Service. Rescinded.

Rule 9.  (1) The repair and service of equipment provided by the department for evidential and preliminary breath alcohol analysis must be at the expense of the agency using the equipment.

(2) The department or the agency using the equipment may arrange to have this service provided either by an appropriate class operator who has been certified in accordance with R 325.2658 or a manufacturer-trained authorized representative approved by the department.

(3) After repair or service and before being placed in service, evidential and preliminary breath alcohol test instruments must be verified for accuracy in accordance with R 325.2653.  Records of verification must be maintained as required by the department.